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External Collector Centrifuge - Initial Qualification

Company Name: _________________		Centrifuge Model: _______________
Serial #: ________________________		
	Installation Qualifications
	Date
	Initials

	[bookmark: _Hlk71183998]Develop Maintenance Plan According to Vendor Requirements (user manual)
	
	

	RPM* set on centrifuge to achieve appropriate RCF** 
Refer to Eastern Health document 14711 Centrifuge RCF (g Force) Determination
	
	

	Check operation & perform speed and timer accuracy check.
	
	

	Check over unit for any physical or mechanical damage.
	
	

	Check condition of power cord & strain relief.
	
	

	Ensure correct operation of safety lid & latch.
	
	

	Ensure correct operation of emergency override if applicable.
	
	

	Check temperature accuracy if applicable
	
	



Initial Calibrations
	Speed (RPM)
	





	Tachometer Calibration 
Timer (min)

	Set
	Display
	
	Set
	Actual

	5000
	
	
	
	

	10000
	
	
	
	

	15000
	
	
	
	


	

*RPM – Revolutions per minute – rpm must be set to achieve proper RCF.
** RCF – Relative centrifugal force or gravitational force. 
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